# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

# Contract title: Supply and installation of Pneumatic Tube System within LAB-OP project, Bar Montenegro

Publication reference: 12633

Columns 1-2 should be completed by the contracting authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the contracting authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words ‘compliant’ or ‘yes’ are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item number | 2. Specifications required | 3. Specifications offered | 4. Notes, remarks, ref to documentation | 5. Evaluation committee’s notes |
| --- | --- | --- | --- | --- |
| Zone Controller and Energy Supply |  |  |  |  |
| 1 | Central.Control.Unit Linux i5-7200 or equivalent. Industrial grade controller unit, fan -less, small and robust, low energy consumption, i5-Glada processor or equivalent, 4GB RAM, 240 GB SSD, 1 Ethernet, 4 USB Displayport. Operating system Linux, Virtualisation of controller (Virtual Box, VMWare...), Windows 7/10 Client, Control unlimited numbers of stations and zones, SQL Database, Client-Server Technology; Features: Multi-Client-Connection, Statistics, User identification, Secure dispatch, Carrier Maintenance Counter, Multiple Fire alarm settings. Reliable industry standard RS 485 databus communication, Linux and virtual applications, Graphical user interface, Data stored in SQL base, Sending and receiving history with indication of destination and date and time.; 1; pcs. Software Add-On; 1; pcs. Gateway 1xin/2xout WM-housing or equivalent; 2; pcs. SWITCHING POWER PACK 24 VDC/5A METAL- box; 2; pcs. TRANSPONDER READER+PROGRAM-UNIT RF12; 1; pcs. UPS POWER SUPPLY 400V 230V 50/60Hz; 1; pcs. RFID CARD TEHNOLOGY MIFARE 1K CLASIK; 150; pcs. EMERGENCYDONGLE 20d +3000 Run TIME or equivalent; 1; pcs. |  |  |  |
| Stations and Diverters |  |  |  |  |
| 2 | TIITAN - STATION 2R2 OD160I/O TI or equivalent- Titan station is used for sending and receiving of transport carriers and it could be installed as end station or middle station, where is the tube system installed in a sequential system. In this option has the station possibility of two entrance (input)/exit (output)-below and above. Sliding plexi-glass door enables safe use and modern design. When entering the station address, to which we would like to send the carrier, the sliding door opens and we put in the carrier, the station automatically detect it with sensors and the station automatically sends the carrier to the destination, where we wish. Upon receipt the carrier falls lightly and quietly into the basket below the station, where it stays till the user pick it out.; 5; pcs. RECEIVING BASKET LATTICE OD 160 RAL 7035; 5; pcs. Lever converter 2-h.TTL to RS232 UniCon; 5; pcs. CARIER-STOP GREY OD160 ARRIVABOVE OLU; 2; pcs. AIR VALVE OD 160 D160,stick, complete; 2; pcs. Carrier-stop grey OD 160 ArrivBelow oLu; 1; pcs. EWS STATION 2R2 OD160 I/O TK1 or equivalent - THe most typical and used EWS station is dedicated sending and receiving of carriers. It is meant as the end station in the PTS. It has a connection to the pipeline on the top, through which it is receiving and transmitting carriers, from below it has an output for receiving and sending of containers and a clipboard for sending the carriers. Under the station we install the receiving basket in which the carriers are gently stopped. Robustness and a small number of moving parts contribute to its typical use. Membrane keyboard with easy to read monitor and signaling LED indicators is user friendly designed.; 5; pcs. RECEIVING BASKET LATTICE OD 160 RAL 7035; 4; pcs. Guide rail Basic Set OD160 R=650; 2; pcs. MultiloadStat.3f OD160 or equivalent - Multiload station is designed to send three carriers in a row. It has three inserts that allow the user to be transmitted to the station three carriers,which are then distributed sequentially and automatically to the desired destinations. Ergonomic membrane keyboard with monitor is user friendly for user and it is protected from dust and dirt. LED indicators displays the status of sending and user can read it also on monitor. The station allows each clippboard to be programmed according to the contents in the carrier; the speed of sending or priorities. As the station is only sending an additional receiving station with a sledge is recommended. (laboratories).; 2; pcs. Guide rail ExtenSet1 OD160-OD200 L=1000; 2; pcs. Slide gate RFI2 OD 160. Slide gate is typically used forreceiving carriers, tube aproaching from the top. For soft arrivals at Open-End-Stations. Divides the blower connection in AC2U Systems. Also included in Slow-Speed-Transportation - Devices.; 2; pcs. Lever converter 2-h.TTL to RS232 UniCon; 2; pcs. REFITING KIT RFI4 TITAN MIFARE; 5; pcs. REFITING KIT RF14 METDOORMIFARE; 2; pcs. BAR CODE SCANER HAND METAPACE USB; 10; pcs. CARIER RACK WIRE 5-FOLD OD160 GREY; 19; pcs. DIVERTER 2-WAY OD160 I/0; 7; pcs. Tube switch optical OD 160 COMPL.W.CAP.; 7; pcs. |  |  |  |
| Blower |  |  |  |  |
| 3 | Blower SD7 w. valve 400V 3,0KW 6,0A 50Hz; 2; pcs. Blower control 6,3-10 3-way-val SD7; 2; pcs. Conecting - material Air for SD7-SD80; 2; pcs. Connecting-material Blower SD 7-SD90; 2; pcs. $SILENCER OD90 FOR BLOWER SD7-SD90; 4; pcs. REDUCTION RING PVC OD160/OD110 H15MM; 2; pcs. AIRTUBE FLEXIBLE OD90/OD110 GREY L=1M; 2; m. REDUCTION PVC GREY OD110/OD90 L135MM; 2; pcs. L-FITING OD160 TUBE/TUBE S=350 MM; 4; pcs. AIR-REDUCTION OD160 CARRIER STOP; 2; pcs. Slow speed device OD110; 2; pcs. |  |  |  |
| Tubing and Installation |  |  |  |  |
| 4 | DISPATCH TUBE 160X3,2 OD160 PVC GREY RAL 7000; 360; m. DISP.BEND 160X3,2 OD160 R800 PVC GREY; 75; pcs. CONNECTING SLEEVE OD160 PVC L150SHORT D167XS3,2MM; 310; pcs. DISP. TUBE 160X3,2 PVC OD160- L5000 TRANSP.; 50; m. DISP.BEND 160X3,2 OD160 R800 PVC TRANSP; 15; pcs. SLEEVE OD 160 TRANSPARENT; 32; pcs. FIRE PROTECT SLEEVE OD160 INCORP/SU.MOU; 30; pcs. Steel sleeve OD 160 L100; 40; pcs. TUBE CLAMP BASIC OD160 D157-163 M8/10; 110; pcs. ADHESIVE TANGIT PVC-U can 1kg; 14; kg. CLEANER TANGIT PBC-U/C can 1000ml; 12; liter. Control kable PVC 3X2X0,6 J-Y(st)Y; 460; m. POWER CABLE 100M PVC 3X2,5QMM NYM-J; 80; m. Carrier-KD-1A-MA-KL 160/134/116-350 1x swiveling cap LEAK-PROOF; 20; pcs. Laboratory bag Vinyl OD160 W330xH300mm 6f; 20; pcs. SAMPLE CARRIER INSERT; 20; pcs. CONSTRUCTION AND CRAFT WORKS- penetrations fi 180 different thicknesses and slabs; 12; compl.. |  |  |  |
| Other |  |  |  |  |
| 5 | SYSTEM INSTALLATION - INSTALLATION OF PIPELINE INSTALATION; 1; compl.. INSTALLATION OF STATIONS AND SWITCHES; 1; compl.. ELECTRONICS INSTALLATION; 1; compl.. PROGRAMMING; 1; compl.. TESTING; 1; compl.. HANDOVER; 1; compl.. PID PROJECT; 1; compl.. Transporting and forwarding; 1; compl.. |  |  |  |
| Spare parts |  |  |  |  |
| 6 | Limit (final) switch-AC optical 281050/CZT6-P0200501; 5; pcs. Relay T-Matic; 2; pcs. Contactor 5,5A; 2; pcs. Power supply 24VDC*5A; 1; pcs. TRANSPORT CARRIER LODGE FI 160 AC; 1; pcs. Control panel of the station type GCT 11/A; 1; pcs. SEALING TAPE FOR CARRIER fi 160; 80; pcs. Cevna objemka PLUS W2 with auger - GIANT 150-170 mm; 20; pcs. Sealing ring-under the hedgehog fi110; 20; pcs. The foam for the basket - big; 10; pcs. |  |  |  |
| Service maintenance for 2 years |  |  |  |  |
| 7 | Service maintanace; 8; pcs. |  |  |  |